## Readme
Draft for review, not to cite or circulate
2019-01-03
Notes:
LBNL is developing a prototype energy model to support the evaluation of energy efficiency for laboratory buildings.
This is the initial version of the model specification, intended for review and obtaining feedback from stakeholders.
Team contact:
Tianzhen Hong: thong@lbl.gov
Kaiyu Sun: ksun@lbl.gov
Sang Hoon Lee: sanghlee@lbl.gov

## Building Description
Prototype Laboratory Modeling Specifications
Lawrence Berkeley National Laboratory, updated on 12-07-2018
| Item | Descriptions | Data Source
Program
| Vintage | N/A
| Location | All 90.1 climate zones | Selection of representative climates based on Briggs et al. (2003)
| Available fuel types | Gas, electricity
| Building Type (Principal Building Function) | Laboratory
| Building Prototype | Laboratory
Form
| Total Floor Area (sq feet) | 90,000 ( 200ft x 150ft x 3 stories) | - Report "Ultra-Low Temperature Freezers: Opening the Door to Energy Savings in Laboratories, ET14PGE1721" - EIA CBECS 2012
| Building shape | Rectangular
| Aspect Ratio | 1.33
| Number of Floors | 3
| Window Fraction (Window-to-Wall Ratio) | 19% (Window Dimensions: 8 windows, 12.5 ft x 6 ft on the long side of facade 6 windows, 12.5 ft x 6 ft on the short side of the façade) | Calculated based on CBECS 2012
| Window Locations | Evenly distributed along four façades
| Shading Geometry | None
| Azimuth | Long side faces South and North
| Thermal Zoning | Three core zones and four perimeter zone, 25-ft wide perimeter zones
Core zones: Laboratory space 15,000 ft2 (50% per floor area), including: (1) lab space with fume hood, 3,000 ft2 (20% lab area); (2) lab equipment corridor, 1,500 ft2 (10% lab area); (3) open lab space, 10,500 ft2 (70% lab area) Perimeter zones: office space 15,000 ft2 (50% per floor area)
| Floor to floor height (feet) | 15
| Floor to ceiling height (feet) | 10
| Glazing sill height (feet) | 2 ft (top of the window is 8 ft high with 6 ft high glass)
Architecture
Exterior walls
| Construction | Mass (pre-cast concrete panel): 8 in. Heavy-Weight Concrete + Wall Insulation + 0.5 in. gypsum board | Construction type: PNNL's CBECS Study
| U-factor (Btu / h * ft2 * °F) and/or R-value (h * ft2 * °F / Btu) | Requirements in codes or standards Nonresidential; Walls, Above-Grade, Steel-Framed Wall U-factor (Btu/h*ft2*°F): ASHRAE 90.1 2004: 0.084, 2007: 0.064, 2010: 0.064, 2013: 0.055, and 2016: 0.055 | Applicable codes or standards (ASHRAE 90.1)
| Dimensions | Based on floor area and aspect ratio
| Tilts and orientations | Vertical
Roof
| Construction | Built-up roof: Roof membrane+roof insulation+metal decking | Construction type: 2003 CBECS Data and PNNL's CBECS Study 2007. Base assembly from 90.1 Appendix A.
| U-factor (Btu / h * ft2 * °F) and/or R-value (h * ft2 * °F / Btu) | Requirements in codes or standards Nonresidential; Roofs, Insulation entirely above deck Roof U-factor (Btu/h*ft2*°F): ASHRAE 90.1 2004: 0.063, 2007: 0.048, 2010:, 0.048, 2013: 0.032, and 2016: 0.032 | Applicable codes or standards (ASHRAE 90.1)
| Dimensions | Based on floor area and aspect ratio
| Tilts and orientations | Horizontal
Window
| Dimensions | Based on window fraction, location, glazing sill height, floor area and aspect ratio
| Glass-Type and frame | Hypothetical window with weighted U-factor and SHGC
| U-factor (Btu / h * ft2 * °F) | Requirements in codes or standards Nonresidential; vertical glazing | Applicable codes or standards (ASHRAE 90.1)
SHGC (all)
| Visible transmittance | Same as above requirements
| Operable area | 0% | Ducker Fenestration Market Data provided by the 90.1 Envelope Subcommittee
Skylight
| Dimensions | No skylights
| Glass-Type and frame | NA
U-factor (Btu / h * ft2 * °F)
SHGC (all)
Visible transmittance
Foundation
| Foundation Type | Slab-on-grade floors (unheated)
| Construction | 8" concrete slab poured directly on to the earth
| Thermal properties for ground level floor U-factor (Btu / h * ft2 * °F) and/or R-value (h * ft2 * °F / Btu) | Requirements in codes or standards Nonresidential; slab-on-grade floors, unheated | Applicable codes or standards (ASHRAE 90.1)
| Thermal properties for basement walls | NA
| Dimensions | Based on floor area and aspect ratio
Interior Partitions
| Construction | 2 x 4 uninsulated stud wall
| Dimensions | Based on floor plan and floor-to-floor height
| Internal Mass | 6 inches standard wood (16.6 lb/ft²)
Air Barrier System
| Infiltration (cfm/ft2 exterior wall) | Above grade exterior wall surface area, adjusted by wind (when fans turn off) Off Peak: 25% of peak infiltration rate (when fans turn on) Perimeter zones Infiltration (cfm/ft2 exterior wall) peak value: ASHRAE 90.1 2004: 0.202, 2007: 0.202, 2010: 0.112, 2013: 0.112, 2016: 0.112 No infiltration for core zones | PNNL-18898. Infiltration Modeling Guidelines for Commercial Building Energy Analysis. Modeled peak infiltration rate may be different for different codes or standards because of their continuous air barrier requirements.
HVAC
System Type
| Heating type | Two Gas boilers | Number of boilers will be determined after sizing simulation. Minimum of 2 for redundancy
| Cooling type | Packaged VAV
| Distribution and terminal units | Laboratory zones: One AHU system and electric stream humidifiers, 100% outdoor air VAV with hot water reheating Non-laboratory zones: One AHU system with VAV with hot water reheating
HVAC Sizing
| Air Conditioning | Autosized to design day
| Heating | Autosized to design day
HVAC Efficiency
| Air Conditioning | Efficiency varies by design cooling capacity Requirements in codes or standards | Applicable codes or standards (ASHRAE 90.1)
| Heating | Boiler efficiency varies by climate location and design heating capacity Requirements in codes or standards
HVAC Control
| Thermostat Setpoint | Laboratory zones: 72°F Cooling/72°F Heating Office zones: 75°F Cooling/70°F Heating | For laboratory spaces: - LAB21 EnergyPlus Benchmarking Model - National Institutes of Health, Design Requirements Manual, Table 6.1.9.1 For office spaces: DOE Office prototype models
| Thermostat Setback | Laboratory zones: no setback Office zones: 80°F Cooling/60°F Heating | For laboratory spaces: LAB21 EnergyPlus Benchmarking Model For office spaces: DOE Office prototype models
| Supply air temperature | Non-lab zones: The usual supply-air-to-room-air temperature difference of 20 deg F (Supply air temperature: Heating 104F, cooling 55F) Lab zones: Supply-airto-room-air temperature difference of 17 deg F (Supply air temperature: Heating 89F, cooling 55F) | -Laboratory Modeling Guideline using ASHRAE 90.1-2007 Appendix G -Report "Ultra-Low Temperature Freezers: Opening the Door to Energy Savings in Laboratories, ET14PGE1721" -ASHRAE 90.1-2016 Exception to G3.1.2.8.1
| Supply air humidity | NA | LAB21 EnergyPlus Benchmarking Model
| Chilled water supply temperatures | NA
| Hot water supply temperatures | 180 F
| Economizers | Non-laboratory zones: Requirements in codes or standards | Applicable codes or standards (ASHRAE 90.1)
| Ventilation | For open lab and equipment corridor spaces: Use the greater of (1) exhaust air flow rate and (2) meeting the cooling load For fume hood lab spaces: follow exhaust air flow rate (15 ACH) For office spaces: DOE office prototype models (See under Outdoor Air) | -For laboratory spaces: ASHRAE Ventilation Standard 62-1999, ASHRAE 62.1-2004, 2007, 2010, 2013, and 2016, Table 6-1Education Facilities Science Laboratories -For office spaces: DOE Office prototype models
| Demand Control Ventilation | No
| Energy Recovery | For laboratory spaces: No For office spaces: Yes
Supply Fan
| Fan schedules | See under Schedules
| Supply Fan Total Efficiency (%) | Depending on the fan motor size and requirements in codes or standards | Requirements in applicable codes or standards for motor efficiency and fan power limitation
| Supply Fan Pressure Drop | Various depending on the fan supply air cfm
Exhaust Fan
| Fan schedules | For laboratory zones: See under Schedules
| Exhaust airflow rate | For open lab and equipment corridor spaces: Use the greatest of (1) 1cfm/ft2 ASHRAE 62.1, (2) 6ACH exhaust air, (3) 90% of the supply fan airflow rate Fume hood: 15 ACH During unoccupied periods reduce to the largest of 50% of zone peak airflow, the minimum outdoor airflow rate, or the airflow rate required to comply with applicable codes or accreditation standards (For 2010, 2013 and 2016, not 2004 and 2007) | -Minimum requirement: A Guide to Navigating Building and Fire Codes for Laboratories (ASHRAE Ventilation Standard 62.1 Educational science laboratories does not apply to the majority of labs.) -Laboratory Modeling Guideline using ASHRAE 90.1-2007 Appendix G - ASHRAE 90.1 6.5.7.2 Laboratory Exhaust Systems - Exception to G3.1.3.13 in 90.1-2016
| Fan Total Efficiency (%) | Depending on the fan motor size and requirements in codes or standards | Requirements in applicable codes or standards for motor efficiency and fan power limitation
| Fan Pressure Drop | Various depending on the fan supply air cfm
Pump
| Pump Type | Service hot water (SWH): constant speed; hot water (HW) pump: variable speed | ASHRAE 90.1 Appendix G
| Rated Pump Head | Pump power assumptions as specified in ASHRAE 90.1 Appendix G | ASHRAE 90.1 Appendix G
| Pump Power | Autosized
Cooling Tower
| Cooling Tower Type | NA
| Cooling Tower Power | NA
Service Water Heating
| SWH type | One main water heater with storage tank
| Fuel type | Natural gas
| Thermal efficiency (%) | Requirements in codes or standards | Applicable codes or standards
| Tank Volume (gal) | 300 | Tank size is based on ASHRAE 90.1 Office large reference model
| Water temperature setpoint | 140 F | PNNL-23269 Enhancements to ASHRAE Standard 90.1 Prototype Building Models
| Water peak flow rate (ft2/min) | Lab zones: 0.00447, Non-lab zones: 0.00758 | Lab zone: ASHRAE 90.1 Hospital ref model lab zone Non-lab zone: ASHRAE 90.1 Office medium
| Water consumption | See under Schedules
Internal Loads & Schedules
Lighting
| Average power density (W/ft2) | See Zone Summary Requirements in codes or standards | ASHRAE 90.1 Table 9.5.1 Lighting Power Densities using the Building Area Method
| Schedule | See under Schedules | Laboratory Modeling Guideline using ASHRAE 90.1-2007 Appendix G
| Daylighting Controls | Requirements in codes or standards | Applicable codes or standards
| Occupancy Sensors | Requirements in codes or standards | Applicable codes or standards
Plug load
| Average power density (W/ft2) | See under Zone Summary (Lab space with fume hood and open lab space: 4.0 W/ft2, lab equipment corridor: 15 W/ft2 Non-laboratory zones: 0.75 W/ft2) | User's Manual for ASHRAE Standard 90.1-2004 (Appendix G) LAB21 EnergyPlus Benchmarking Model Lab equipment corridor power density based on industry input
| Schedule | See under Schedules | Laboratory Modeling Guideline using ASHRAE 90.1-2007 Appendix G
Occupancy
| Average people | See under Zone Summary (Lab space with fume hood and open lab space: 200 ft2/person, lab equipment corridor: none, Non-laboratory zones: 200 ft2/person) | a. ASHRAE 62.1 Table 6.1 Educational Facilities: Sciene laboratories 40 ft2/person b. Hospital reference model: LAB zone 200 ft2/person c. Report "Ultra-Low Temperature Freezers: Opening the Door to Energy Savings in Laboratories, ET14PGE1721": 200 ft2/person d. Lab21 Laboratory space occupancy: 500 ft2/person
| Schedule | See under Schedules | Laboratory Modeling Guideline using ASHRAE 90.1-2007 Appendix G
Misc.
Elevator
| Quantity | 2 | DOE Commercial Reference Building Models of the National Building Stock
| Motor type | Hydraulic
| Peak Motor Power Watts per elevator | 16055
| Heat Gain to Building | Interior
| Peak Fan/lights Power Watts per elevator | 161.9 | 90.1 Mechanical Subcommittee, Elevator Working Group
| Motor and fan/lights Schedules | See under Schedules | DOE Commercial Reference Building TSD and models (V1.3_5.0) and Addendum DF to 90.1-2007
Exterior Lighting
| Peak Power | Based on design assumptions for façade, parking lot, entrance, etc. and requirements in codes or standards | Applicable codes or standards
| Schedule | See under Schedules and control requirements in codes or standards | Applicable codes or standards
References
1 | UC Berkeley and LBNL. 2009. LAB21 EnergyPlus Benchmarking Model
2 | PG&E, Southern California Edison, San Diego Gas and Electric Company. 2016. Ultra-Low Temperature Freezers: Opening the Door to Energy Savings in Laboratories, ET14PGE1721
3 | EIA, 2012, Commcer Buildings Energy Consumption Survey (CBECS) 2012
4 | ASHRAE 90.1 2004, 2007, 2010,2013, 2016, Energy Standard for Buildings Except Low-Rise Residential Buildings
5 | ASHRAE 62.1 2004, 2007, 2010, 2013 Ventilation for Acceptable Indoor Air Quality
6 | ASHRAE 62-1999 Ventilation for Acceptable Indoor Air Quality
7 | Labs21. 2008, Laboratory Modeling Guideline using ASHRAE 90.1-2007 Appendix G
8 | Briggs, R.S., R.G. Lucas, and Z.T. Taylor. 2003. Climate Classification for Building Energy Codes and Standards: Part 2—Zone Definitions, Maps, and Comparisons. ASHRAE Transactions 109(2).
9 | PNNL's CBECS Study. 2007. Analysis of Building Envelope Construction in 2003 CBECS Buildings. Dave Winiarski, Mark Halverson, and Wei Jiang. Pacific Northwest National Laboratory. March 2007.
10 | PNNL's CBECS Study. 2006. Review of Pre- and Post-1980 Buildings in CBECS – HVAC Equipment. Dave Winiarski, Wei Jiang and Mark Halverson. Pacific Northwest National Laboratory. December 2006.
11 | Gowri K, DW Winiarski, and RE Jarnagin. 2009. Infiltration modeling guidelines for commercial building energy analysis . PNNL-18898, Pacific Northwest National Laboratory, Richland, WA. http://www.pnl.gov/main/publications/external/technical_reports/PNNL-18898.pdf
12 | Goel S, M Rosenberg, R Athalye, Y Xie, W Wang, R Hart, J Zhang, V Mendon. 2014. Enhancements to ASHRAE Standard 90.1 Prototype Building Models. PNNL-23269, Pacific Northwest National Laboratory, Richland, Washington. http://www.pnnl.gov/main/publications/external/technical_reports/PNNL-23269.pdf
13 | P Mathew, etc. Advanced Benchmarking for Complex Building Types: Laboratories as an Exemplar. ACEEE conference 2010.
14 | DOE. A Guide to Navigating Building and Fire Codes for Laboratories.
15 | National Institutes of Health. Design Requirements Manual. Issued on 12/12/2016, Rev 4/24/2019

## Zone Summary
Zone Summary
[7 empty] | ASHRAE 90.1 2004 | ASHRAE 90.1 2007 | ASHRAE 90.1 2010 | ASHRAE 90.1 2013 | ASHRAE 90.1 2016
Zone | Area [ft²] | Conditioned [Y/N] | Volume [ft³] | Multipliers | Gross Wall Area [ft²] | Window Glass Area [ft²] | Lighting 2 [W/ft²] |  |  |  |  | People 4 [ft²/person] | Number of People | Plug and Process3 [W/ft²]
CORE_BOTTOM_Open_Lab | 10500 | Yes | 105000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | 200 | 52.5 | 4
CORE_BOTTOM_FumeHood_Lab | 3000 | Yes | 30000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | 200 | 15 | 4
CORE_BOTTOM_Lab_Equipment_Corridor | 1500 | Yes | 15000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | - | 0 | 15
CORE_MID_Open_Lab | 10500 | Yes | 105000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | 200 | 52.5 | 4
CORE_MID_FumeHood_Lab | 3000 | Yes | 30000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | 200 | 15 | 4
CORE_MID_Lab_Equipment_Corridor | 1500 | Yes | 15000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | - | 0 | 15
CORE_TOP_Open_Lab | 10500 | Yes | 105000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | 200 | 52.5 | 4
CORE_TOP_FumeHood_Lab | 3000 | Yes | 30000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | 200 | 15 | 4
CORE_TOP_Lab_Equipment_Corridor | 1500 | Yes | 15000 | 1 | 0 | 0 | 1.4 | 1.4 | 1.81 | 1.81 | 1.45 | - | 0 | 15
PERIMETER_TOP_ZN_3 | 4375 | Yes | 43750 | 1 | 2000 | 900 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 21.875 | 0.75
PERIMETER_TOP_ZN_2 | 3125 | Yes | 31250 | 1 | 1500 | 675 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 15.625 | 0.75
PERIMETER_TOP_ZN_1 | 4375 | Yes | 43750 | 1 | 2000 | 900 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 21.875 | 0.75
PERIMETER_TOP_ZN_4 | 3125 | Yes | 31250 | 1 | 1500 | 675 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 15.625 | 0.75
PERIMETER_BOT_ZN_3 | 4375 | Yes | 43750 | 1 | 2000 | 900 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 21.875 | 0.75
PERIMETER_BOT_ZN_2 | 3125 | Yes | 31250 | 1 | 1500 | 675 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 15.625 | 0.75
PERIMETER_BOT_ZN_1 | 4375 | Yes | 43750 | 1 | 2000 | 900 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 21.875 | 0.75
PERIMETER_BOT_ZN_4 | 3125 | Yes | 31250 | 1 | 1500 | 675 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 15.625 | 0.75
PERIMETER_MID_ZN_3 | 4375 | Yes | 43750 | 1 | 2000 | 900 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 21.875 | 0.75
PERIMETER_MID_ZN_2 | 3125 | Yes | 31250 | 1 | 1500 | 675 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 15.625 | 0.75
PERIMETER_MID_ZN_1 | 4375 | Yes | 43750 | 1 | 2000 | 900 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 21.875 | 0.75
PERIMETER_MID_ZN_4 | 3125 | Yes | 31250 | 1 | 1500 | 675 | 1 | 1 | 0.9 | 0.82 | 0.79 | 200 | 15.625 | 0.75
TOPFLOOR_PLENUM 5 | 15000 | No | 150000 | 1 | 3500 | 0 | - | - | - | - | - | - | 0
MIDFLOOR_PLENUM | 15000 | No | 150000 | 1 | 3500 | 0 | - | - | - | - | - | - | 0
FIRSTFLOOR_PLENUM | 15000 | No | 150000 | 1 | 3500 | 0 | - | - | - | - | - | - | 0
TOTAL1 | 90000 |  | 1350000 |  | 31500 | 9450 | [6 empty] | 427.5
AREA WEIGHTED AVERAGE | [6 empty] | 1.2 | 1.2 | 1.355 | 1.315 | 1.12 | 190 |  | 2.925
1. Only volume, and gross wall area include unconditioned space.
2. Listed lighting power density is based on applicable requirements in ASHRAE Standard 90.1 Table 9.5.1 for office and Table 9.6.1 for laboratory.
3. Laboratry zones electric power density, 4 W/ft2 is basd on the Lab21 benchmark model and hospital reference model lab zones
4. Laboarty occupant densities are based on the following references:
a. ASHRAE 62.1 Table 6.1 Educational Facilities: Sciene laboratories 40 ft2/person
b. Hospital reference model: LAB zone 200 ft2/person
c. ULTF occupancy: 200 ft2/person
d. Lab21 Laboratory space occupancy: 500 ft2/person
5. Plenum is only for office zones
We may want to use 200ft2/person. As 40 ft2/persion by ASHRAE 62.1 is based on the educational facility laboratory, the normal laboratory people density may be less than the educational purpose.

## Outdoor Air
Minimum Outdoor Ventilation Air Requirements
|  |  |  | Total Occupants | Total OSA Ventilation (cfm/zone)1,2 |  |  |  |  | Total OSA Ventilation (cfm/ft2)
Zone | Area (ft2) | Multipliers | Assumed Space Type | 62.1-2004, 2007, 2010, 2013, 2016 | 90.1-2004 | 90.1-2007 | 90.1-2010 | 90.1-2013 | 90.1-2016 | 90.1-2004 | 90.1-2007 | 90.1-2010 | 90.1-2013 | 90.1-2016
CORE_BOTTOM_Open_Lab | 10500 | 1 | Open Laboratory | 52.5 | 10500 | 10500 | 10500 | 10500 | 10500 | 1 | 1 | 1 | 1 | 1
CORE_BOTTOM_FumeHood_Lab | 3000 | 1 | Lab with fume hood | 15 | 7500 | 7500 | 7500 | 7500 | 7500 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5
CORE_BOTTOM_Lab_Equipment_Corridor | 1500 | 1 | Lab equipment corridor | 0 | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD
CORE_MID_Open_Lab | 10500 | 1 | Open Laboratory | 52.5 | 10500 | 10500 | 10500 | 10500 | 10500 | 1 | 1 | 1 | 1 | 1
CORE_MID_FumeHood_Lab | 3000 | 1 | Lab with fume hood | 15 | 7500 | 7500 | 7500 | 7500 | 7500 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5
CORE_MID_Lab_Equipment_Corridor | 1500 | 1 | Lab equipment corridor | 0 | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD
CORE_TOP_Open_Lab | 10500 | 1 | Open Laboratory | 52.5 | 10500 | 10500 | 10500 | 10500 | 10500 | 1 | 1 | 1 | 1 | 1
CORE_TOP_FumeHood_Lab | 3000 | 1 | Lab with fume hood | 15 | 7500 | 7500 | 7500 | 7500 | 7500 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5
CORE_TOP_Lab_Equipment_Corridor | 1500 | 1 | Lab equipment corridor | 0 | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD | TBD
PERIMETER_TOP_ZN_3 | 4375 | 1 | Office space | 21.875 | 437.5 | 371.875 | 371.875 | 371.875 | 371.875 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_TOP_ZN_2 | 3125 | 1 | Office space | 15.625 | 312.5 | 265.625 | 265.625 | 265.625 | 265.625 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_TOP_ZN_1 | 4375 | 1 | Office space | 21.875 | 437.5 | 371.875 | 371.875 | 371.875 | 371.875 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_TOP_ZN_4 | 3125 | 1 | Office space | 15.625 | 312.5 | 265.625 | 265.625 | 265.625 | 265.625 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_BOT_ZN_3 | 4375 | 1 | Office space | 21.875 | 437.5 | 371.875 | 371.875 | 371.875 | 371.875 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_BOT_ZN_2 | 3125 | 1 | Office space | 15.625 | 312.5 | 265.625 | 265.625 | 265.625 | 265.625 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_BOT_ZN_1 | 4375 | 1 | Office space | 21.875 | 437.5 | 371.875 | 371.875 | 371.875 | 371.875 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_BOT_ZN_4 | 3125 | 1 | Office space | 15.625 | 312.5 | 265.625 | 265.625 | 265.625 | 265.625 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_MID_ZN_3 | 4375 | 1 | Office space | 21.875 | 437.5 | 371.875 | 371.875 | 371.875 | 371.875 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_MID_ZN_2 | 3125 | 1 | Office space | 15.625 | 312.5 | 265.625 | 265.625 | 265.625 | 265.625 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_MID_ZN_1 | 4375 | 1 | Office space | 21.875 | 437.5 | 371.875 | 371.875 | 371.875 | 371.875 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
PERIMETER_MID_ZN_4 | 3125 | 1 | Office space | 15.625 | 312.5 | 265.625 | 265.625 | 265.625 | 265.625 | 0.1 | 0.085 | 0.085 | 0.085 | 0.085
TOTAL | 90000 |  |  | 427.5 | 58500 | 57825 | 57825 | 57825 | 57825 | 0.65 | 0.6425 | 0.6425 | 0.6425 | 0.6425
1. The office ventilation requirements are based on the ASHRAE Standard 62.1.
2. According to ASHRAE 62.1 Table 6-1, the minimum requirement for laboratory ventilation rate is 10 cfm/person + 0.18 cfm/ft2. Table 6-4 provides the minimum exhaust rate, 1 cfm/ft2 for educational science laboratories.
Laboratry outdoor air ratio requirement is 100%. The outdoor airflow requirement uses the exhaust airflow rate 1 cfm/ft2 to avoid contamination.
3. Ventilation for lab space with fume hood is based on industry input

## Schedules
Schedule | Type | Through | Day of Week | 1 am | 2 am | 3 am | 4 am | 5 am | 6 am | 7 am | 8 am | 9 am | 10 am | 11 am | Noon | 1 pm | 2 pm | 3 pm | 4 pm | 5 pm | 6 pm | 7 pm | 8 pm | 9 pm | 10 pm | 11 pm | 12 pm
Internal Loads Schedules
Lab_LIGHT_SCH | Fraction | Through 12/31 | Weekday | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.3 | 0.5 | 0.9 | 0.9 | 0.9 | 0.9 | 0.8 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.5 | 0.5 | 0.3 | 0.3 | 0.2 | 0.2 | 0.2 | Schedule updated to ASHRAE 90.1 Appendix G laboratory modeling guide
|  |  | Sat, Sun, Hol, Other | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.4 | 0.4 | 0.4 | 0.4 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
|  |  | SummerDesign | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lab_EQUIP_SCH | Fraction | Through 12/31 | Weekday | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.3 | 0.4 | 0.5 | 0.5 | 0.5 | 0.4 | 0.5 | 0.5 | 0.5 | 0.5 | 0.4 | 0.3 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | Schedule updated to ASHRAE 90.1 Appendix G laboratory modeling guide
|  |  | Sat, Sun, Hol, Other | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.3 | 0.3 | 0.4 | 0.4 | 0.4 | 0.4 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
|  |  | SummerDesign | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lab_OCC_SCH | Fraction | Through 12/31 | Weekday | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.1 | 0.2 | 0.9 | 0.9 | 0.45 | 0.45 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.3 | 0.1 | 0.1 | 0.1 | 0.05 | 0.05 | 0.05 | Schedule updated to ASHRAE 90.1 Appendix G laboratory modeling guide
|  |  | Sat, Sun, Hol, Other | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.1 | 0.1 | 0.3 | 0.3 | 0.3 | 0.3 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
|  |  | SummerDesign | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lab_OCC_Activity | Watt | Through 12/31 | Constant | 167 | [23 empty] | Ref: https://www.engineeringtoolbox.com/met-metabolic-rate-d_733.html
Office_LIGHT_SCH | Fraction | Through 12/31 | Weekday | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.1 | 0.1 | 0.3 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.5 | 0.3 | 0.3 | 0.2 | 0.2 | 0.1 | 0.05
|  |  | Saturday | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.1 | 0.1 | 0.3 | 0.3 | 0.3 | 0.3 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
|  |  | Sun, Hol, Other | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
|  |  | SummerDesign | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Office_EQUIP_SCH | Fraction | Through 12/31 | Weekday | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.9 | 0.9 | 0.9 | 0.9 | 0.8 | 0.9 | 0.9 | 0.9 | 0.9 | 0.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
|  |  | Saturday | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5 | 0.5 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
|  |  | Sun, Hol, Other | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
|  |  | SummerDesign | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Office_OCC_SCH | Fraction | Through 12/31 | Weekday | 0 | 0 | 0 | 0 | 0 | 0 | 0.1 | 0.2 | 0.95 | 0.95 | 0.95 | 0.95 | 0.5 | 0.95 | 0.95 | 0.95 | 0.95 | 0.3 | 0.1 | 0.1 | 0.1 | 0.1 | 0.05 | 0.05
|  |  | Saturday | 0 | 0 | 0 | 0 | 0 | 0 | 0.1 | 0.1 | 0.3 | 0.3 | 0.3 | 0.3 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.05 | 0.05 | 0 | 0 | 0 | 0 | 0
|  |  | Sun, Hol, Other | 0 | 0 | 0 | 0 | 0 | 0 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0
|  |  | SummerDesign | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BLDG_ELEVATORS | Fraction | Through 12/31 | Weekday | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.35 | 0.69 | 0.43 | 0.37 | 0.43 | 0.58 | 0.48 | 0.37 | 0.37 | 0.46 | 0.62 | 0.12 | 0.04 | 0.04 | 0 | 0 | 0
|  |  | Sat, Sun, Hol, Other | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.16 | 0.14 | 0.21 | 0.18 | 0.25 | 0.21 | 0.13 | 0.08 | 0.04 | 0.05 | 0.06 | 0 | 0 | 0 | 0 | 0 | 0
Exterior Lighting Schedule
BLDG_EXTERIOR_LIGHT | Fraction | Through 12/31 | All Days | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
(AstronomicalClock control)
Service Water Heater Load Schedule
BLDG_SWH_SCH | Fraction | Through 12/31 | Weekday, SummerDesign | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.07 | 0.19 | 0.35 | 0.38 | 0.39 | 0.47 | 0.57 | 0.54 | 0.34 | 0.33 | 0.44 | 0.26 | 0.21 | 0.15 | 0.17 | 0.08 | 0.05 | 0.05 | The same as office medium, but minimum changed to 0.05 for night hours considering lab zone occupant schedule
|  |  | Sat, Sun, Hol, Other, WinterDesign | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.07 | 0.11 | 0.15 | 0.21 | 0.19 | 0.23 | 0.2 | 0.19 | 0.15 | 0.13 | 0.14 | 0.07 | 0.07 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
Infiltration Schedule
INFIL_SCH_PNNL | fraction | Through 12/31 | WD, SummerDesign | 1 | 1 | 1 | 1 | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 1 | 1
|  |  | Sat, WinterDesign | 1 | 1 | 1 | 1 | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | Sun, Hol, Other | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
HVAC Schedules
FumeHood_Schedule | fraction | Through 12/31 | Weekday, SummerDesign, WinterDesign | 0.53 | 0.53 | 0.53 | 0.53 | 0.53 | 0.53 | 0.53 | 0.55 | 0.6 | 0.98 | 0.98 | 0.98 | 0.73 | 0.98 | 0.98 | 0.98 | 0.98 | 0.65 | 0.55 | 0.55 | 0.53 | 0.53 | 0.53 | 0.53
|  |  | Weekend, Hol | 0.53 | 0.53 | 0.53 | 0.53 | 0.53 | 0.53 | 0.53 | 0.55 | 0.55 | 0.65 | 0.65 | 0.65 | 0.65 | 0.55 | 0.55 | 0.55 | 0.55 | 0.55 | 0.53 | 0.53 | 0.53 | 0.53 | 0.53 | 0.53
Lab_HVACOperationSchd | On/off | Through 12/31 | All | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
(Fan Schedule)
Office_HVACOperationSchd | On/off | Through 12/31 | Weekday, SummerDesign | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0
(Fan Schedule) |  |  | Saturday, WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
|  |  | Sun, Hol, Other | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lab_HTGSETP_SCH | Temperature | Through 12/31 | Weekday, WinterDesign | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72
| (°F) |  | Sat, Sun, Hol, Other | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72
Lab_CLGSETP_SCH | Temperature | Through 12/31 | Weekday, SummerDesign | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72
| (°F) |  | Sat, Sun, Hol, Other | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72 | 72
Office_HTGSETP_SCH | Temperature | Through 12/31 | Weekday | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 64.04 | 68 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 60.08 | 60.08
| (°F) |  | Saturday | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 64.04 | 68 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08
|  |  | Sun, Hol, Other | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 60.08
|  |  | WinterDesign | 60.08 | 60.08 | 60.08 | 60.08 | 60.08 | 63.68 | 67.28 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 69.8 | 60.08 | 60.08
Office_CLGSETP_SCH | Temperature | Through 12/31 | Weekday | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 78.08 | 77 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 80.06 | 80.06
| (°F) |  | Saturday | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 78.26 | 77 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 80.06 | 80.06
|  |  | Sun, Hol, Other | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 80.06
|  |  | SummerDesign | 80.06 | 80.06 | 80.06 | 80.06 | 80.06 | 78.26 | 77 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 75.2 | 80.06 | 80.06
CW-Loop-Temp-Schedule | Temperature (°F) | Through 12/31 | All | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06 | 44.06
HW-Loop-Temp-Schedule | Temperature (°F) | Through 12/31 | All | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6 | 179.6
Lab_MinOA_MotorizedDamper_Sched | fraction | Through 12/31 | All | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Office_MinOA_MotorizedDamper_Sched | fraction | Through 12/31 | Weekday, SummerDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0
|  |  | Sat, WinterDesign | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
|  |  | Sun, Hol, Other | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lab Exhaust Fan | on/off | Through 12/31 | All | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
MinOA_Sched | Fraction | Through 12/31 | All | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Notes: | The schedules are also subject to changes in different models based on applicable code requrirements triggered by cllimate zone, system capacity, control type, or other criteria.